# Flow Cell Selection Guide

CRI Sequencing Core Facility.

Last updated: April 21, 2026

## Flow Cell Capacity Table

| Flow cell* | Max reads** |
| --- | --- |
| NextSeq 2000 P4 – 100 Cycles | 1.8B |
| NextSeq 2000 P4 – 200 Cycles | 1.8B |
| NextSeq 2000 P3 – 50 Cycles | 1.2B |
| NextSeq 2000 P3 – 100 Cycles | 1.2B |
| NextSeq 2000 P3 – 200 Cycles | 1.2B |
| NextSeq 2000 P3 – 300 Cycles | 1.2B |
| NextSeq 2000 P2 – 100 Cycles | 400M |
| NextSeq 2000 P2 – 200 Cycles | 400M |
| NextSeq 2000 P2 – 300 Cycles | 400M |
| NextSeq 2000 P1 – 100 Cycles | 100M |
| NextSeq 2000 P1 – 300 Cycles | 100M |
| NextSeq 2000 P1 – 600 Cycles | 100M |
| MiSeq v3 – 150 Cycles | 25M |
| MiSeq v3 – 600 Cycles | 25M |
| MiSeq v2 – 50 Cycles | 15M |
| MiSeq v2 – 300 Cycles | 15M |
| MiSeq Micro v2 – 300 Cycles | 4M |
| MiSeq Nano v2 – 300 Cycles | 1M |

* Cycles: Total number of sequencing cycles (i.e., total bases sequenced per read). For example, a 100‑cycle kit supports SR100 or PE50 (50 cycles per read × 2 reads).

** M = million; B = billion. Read counts are typically listed as single‑read equivalents (SR). In paired‑end (PE) mode, the instrument generates approximately twice the number of reads (two reads per cluster).

When planning throughput, please account for additional reads beyond your sequencing needs. This is because:

- A portion of the sequencing reads will be removed during quality filtering.

- PhiX is routinely spiked into libraries for quality control.

PhiX spike‑in typically ranges from 10% to 50%, depending on library complexity. Low‑complexity libraries (e.g., amplicon‑seq and some ChIP‑seq libraries) generally require higher PhiX percentages to maintain base diversity.

Example: For a PE100 RNA‑seq experiment targeting 25M reads per sample, a P2 200‑cycle flow cell (maximum ~400M reads in SR equivalents) can typically accommodate ~12–14 samples. Although 400/25 = 16 theoretically, additional capacity is reserved to account for PhiX and read filtering.

## Typical experimental requirement (for human and mouse samples)

RNA‑seq

- Read type/length: At least PE50 (longer reads recommended for splicing/isoform analysis)

- Depth: 20–25M reads per sample

ChIP‑seq

- Read type/length: SR or PE; PE is recommended because it can improve library complexity and mapping rate

- Depth:

  - Narrow peaks: >10M reads

  - Broad peaks: >20M reads

ATAC‑seq

- Read type/length: PE50 to PE100; SR is not recommended

- Depth: 25–200M reads, depending on experimental goals (e.g., footprinting analysis requires more reads)

CUT&RUN

- Read type/length: At least PE25; SR is not recommended

- Depth: 3–8M reads

## Additional Note

Samples derived from cancer genomes often benefit from longer reads (and/or paired‑end sequencing) to improve mapping performance, especially in the presence of structural variation, copy-number changes, and repetitive regions.